## Sheet1
| Patient Lift |  | Annual Preventative Maintenance & Performance-verification
|  | Bldg: |  | Room: |  |  |  | Date: / /
Model: |  | Serial #: |  |  | ECN #:
Lift Function |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Up/Down, E.Down, Traverse, On/Off, E. On/Off |  | Press each button on hand control for 5 seconds to ensure each button/function works continuously. If applicable, test emergency lowering. |  |  |  |  | Remote: 2 4 6 button
Limit Switches |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Upper Limit Switch |  | Hold UP until the strap is fully spooled. Motor should stop automatically.
Lower Limit Switch |  | Hold DOWN, and let the strap fully unspool. Motor should stop automatically before reverse winding.
Slack Tape Switch |  | Hold DOWN, and lift up on the carry bar. Motor should stop automatically
Hand Control |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Grommet, Tubing, and Controller |  | Inspect for damage
Mechanical Function |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Load Test per CSA Z10535.2-17 |  | Lift 100% SWL, of the lift, 20" off the ground. If lifting range <5ft, then test the full range |  |  |  |  | Weight used:
Trolley Wheels |  | Inspect wheels for excessive wear and tear
Trolley Failsafe |  | Ensure pin is intact
Motor, Gears, and Traverse Drive |  | Inspect for damage and excessive noise. Add grease.
Carry Bar |  | Inspect for damage
Lift Strap |  | Inspect full length of the strap for fraying and wear.
Electrical Functions |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Charging Block |  | Is the charging block functioning and in good condition?
Battery Inspection |  | Check physical condition and ensure battery has install date less than 3 years old |  |  |  |  | Batt Date:
LCD Screen |  | Check that LCD works properly
LED Light |  | Check LED: green when on, amber when charging
Circuit Board |  | Is it functioning properly? |  |  |  |  | Original Lifts:
Air switches |  | Are they leaking?
Miscellaneous |  | Description |  |  | Pass | Fail | Notes |  | Fixed?
Casing |  | Check for cracks or wear in case
PM Reset |  | Enter program mode; Record # of lifts; Reset the PM Lift Counter |  |  |  |  | Total lifts: PM lifts: Mfg. date:
PM Sticker |  | Complete and place a PM sticker on the lift.
SWL Sticker |  | Ensure SWL information is still accurate and readable
Performed By: |  | /